## Slide 1: Permanent Magnet and Electromagnet Designs for Nonlinear FFA Fields
Enabling fixed-tune non-scaling FFAs

## Slide 2: Application
Dejan Trbojevic proposes to accelerate protons with a rapid cycle (100s of Hz to kHz) for FLASH hadron therapy
  Many machine cycles within 100ms allow different energy (depth) slices to be done as well as being quick enough to trigger the biological FLASH effect
Need an FFA ring which also has fixed tunes
  But needs to fit in the hospital, so smaller than a scaling FFA  nonlinear fixed-tune NS-FFA

## Slide 3: BMAD and Muon1 Magnet Models
BMAD has some sort of zero-length end field model that may or may not be correct
  Particularly not for magnets with many combined high-order multipoles when the FFA beam is off-axis and coming in at a steep angle
Muon1 has a Maxwellian fringe field model with finite length
  But you have to choose the length yourself!
We also changed from sectors to rectangles but that turned out to be a small effect

## Slide 4: Lattice Converted from BMAD
Multipoles dipole…dodeca
Three cells for field wrap
This worked well in BMAD, fixed stable tunes from 10MeV – 250MeV

## Slide 5: Field Profiles

## Slide 6: Original Lattice in Muon1
It’s unstable at the low end no matter what fringe field length I choose
MeV
Cell tune

## Slide 7: cos(cell tunes)
Even with a fringe=1mm it is unstable, although closed orbits exist
MeV

## Slide 8: Peak Fields on Orbit (vs. Energy)
MeV
T

## Slide 9: This cell can be easily fixed using an algorithm I showed last year

## Slide 10: Algorithm
Runge-Kutta 4th order tracking step
Loop to get trajectory in cell
Finite difference to get transfer matrix
  Also gives tunes if orbit is closed
Iterate (Newton) to find closed orbit
Loop over all FFA energies
Finite difference parameters to get response matrix of tune functions to multipole changes
Iterate (optimiser) to find fixed-tune lattice
Muon1
Muon1 cell optics mode
Muon1 FFA optics mode
Fixed-tune FFA optimiser
(parallel)

## Slide 11: Tunes After Correction in Muon1
MeV
Cell tune
Tunes stable from 10MeV – 250MeV and somewhat constant, although could be improved.
(Wiggly near low energy end because rigidity changes fast there).
Qx = 0.1547 ± 0.0013
Qy = 0.0571 ± 0.0034

## Slide 12: Field Profiles After Correction

## Slide 13: Peak Fields After Correction
MeV
T

## Slide 14: Orbits and Fields in Muon1

## Slide 15: Increase Low Energy to 16.5MeV
Use a bigger cyclotron as an injector
  Momentum ratio 5.31×  4.13×, easier to do
Tunes a little better but still wiggly
Qx = 0.1547 ± 0.0008
Qy = 0.0581 ± 0.0031

## Slide 16: Allow 14- and 16-pole Components
That’s noticably better (2-3× less variation)
Probably polynomial isn’t the best field basis
  Use Fourier series transversely instead?
Qx = 0.1545 ± 0.0004
Qy = 0.0577 ± 0.0010

## Slide 17: Field Profiles 12- vs. 16-pole
About 1 part in 1000 change

## Slide 18: Orbits and Fields in Muon1

## Slide 19: Permanent Magnet Designs (F)
cm grid
NdFeB material
N42EH grade
Br=1.3T
107.28cm2 area
64.4 x 16mm aperture
Good field X±26.2mm
1.88T max field

## Slide 20: Permanent Magnet Designs (D)
cm grid
NdFeB material
N42EH grade
Br=1.3T
119.39cm2 area
51.8 x 16mm aperture
Good field X±19.9mm
1.85T max field

## Slide 21: Magnet Field Error Graph (F)
~2×10-4 relative error

## Slide 22: Magnet Field Error Graph (D)
~1×10-4 relative error

## Slide 23: Permanent Magnet Wedges
We bought enough wedges to make two 45mm test sections of the nonlinear D magnet

## Slide 24: Now onto the presentation of the electromagnet (if time)

## Slide 25: Fixed-Tune Non-Scaling FFAs
Design for RAL FETS-FFA (3-12MeV p)
Analysis for cell of two thin lenses

## Slide 26: Introduction
Recent studies by Dejan Trbojevic have confirmed that non-scaling FFAs can have their tune dependence on momentum flattened by adding non-linear components to the magnet fields, although not necessarily for an unlimited momentum range.
Quadrupole affects tune level, sextupole affects dQ/dp, octupole affects d2Q/dp2, etc.

## Slide 27: FETS-FFA requirements
3-12MeV protons
4m average radius
Several 1m drifts (one per cell here)
Fixed tunes: DQx,y<0.01 in ring, <0.001 per cell
Qx,y tunable over full range of 1 in ring
  For exploration of the tune plane in R&D
  I chose 12 cells so this is 1/12 range per cell
Geometrical aperture 1250mm.mrad (large)

## Slide 28: Algorithm
Runge-Kutta 4th order tracking step
Loop to get trajectory in cell
Finite difference to get transfer matrix
  Also gives tunes if orbit is closed
Iterate (Newton) to find closed orbit
Loop over all FFA energies
Finite difference parameters to get response matrix of tune functions to multipole changes
Iterate (optimiser) to find fixed-tune lattice
Muon1
Muon1 cell optics mode
Muon1 FFA optics mode
Fixed-tune FFA optimiser
(parallel)

## Slide 29: Field Model
Midplane field is product of:
  Transverse polynomial up to dodecapole
    Rough initial field coefficients up to octupole were found by Muon1 genetic algorithm before using fixed-tune FFA optimiser
  Longitudinal integral of Gaussian at both ends
    Fringe length given by s=6cm
Off-midplane fields extrapolated by repeated derivatives of this product of three functions

## Slide 30: Fixed-Tune Cell Orbits
Target cell tunes Qx=0.216, Qy=0.213

## Slide 31: Fixed-Tune Cell Fields
F field +0.15 to -0.43T, D field -0.42 to -0.02T

## Slide 32: Cell Tunes
Target cell tunes Qx=0.216, Qy=0.213

## Slide 33: Field Profiles
Only plotted in beam region, including sagitta
3MeV, ring inside 
 12MeV, ring outside

## Slide 34: Field vs. Kinetic Energy

## Slide 35: Field vs. Momentum
Average field should be proportional to momentum
Not exactly because I used “average” position in each magnet

## Slide 36: Gradient vs. Momentum
In a thin-lens model, the gradients should be proportional to momentum
Here they are not, I suspect edge focussing

## Slide 37: Tunes Beyond Energy Range
Target cell tunes Qx=0.216, Qy=0.213
Hits magnet bounding box

## Slide 38: Orbits Beyond Energy Range

## Slide 39: Beta x,y vs. Kinetic Energy

## Slide 40: Tune Range Requirement
Qx,y tunable over full range of 1 in ring
  For exploration of the tune plane in R&D
  I chose 12 cells so this is 1/12 range per cell
Qx=0.216, Qy=0.213 would be ring tune of (2.592,2.556)
So choose 2.5 to 3.5 ring tune range in both planes
  0.2083 to 0.2917 cell tune (all combinations)

## Slide 41: Four Tune “Corners”
Qx=0.2083, Qy=0.2083
Qx=0.2917, Qy=0.2083
Qx=0.2083, Qy=0.2917
Qx=0.2917, Qy=0.2917

## Slide 42: Cell Tunes at Corners
Qx=0.2083, Qy=0.2083
Qx=0.2917, Qy=0.2083
Qx=0.2083, Qy=0.2917
Qx=0.2917, Qy=0.2917

## Slide 43: Cell Tunes at Corners (zoom)
Qx=0.2083, Qy=0.2083
Qx=0.2917, Qy=0.2083
Qx=0.2083, Qy=0.2917
Qx=0.2917, Qy=0.2917

## Slide 44: Field Profiles of Corner Designs
Qx=0.2083, Qy=0.2083
Qx=0.2917, Qy=0.2083
Qx=0.2083, Qy=0.2917
Qx=0.2917, Qy=0.2917

## Slide 45: Designs Summary Table
Cell Qx | Cell Qy | Max tune error | Max field (T) | Orbit range (m)
0.2083 | 0.2083 | 0.00074 | 0.5204 | 0.641
0.2083 | 0.2917 | 0.00185 | 0.6128 | 0.688
0.2917 | 0.2083 | 0.00168 | 0.4466 | 0.389
0.2917 | 0.2917 | 0.00037 | 0.4004 | 0.381
0.216 | 0.213 | 0.00126 | 0.4310 | 0.573
Qx low, Qy high corner is most “unnatural”
  Largest orbit excursion, highest field
    Had to let optimiser vary dipole to stay in magnet box
Higher Qx tunes greatly improve orbit range

## Slide 46: Half-way Interpolation Result
What happens if you average the fields (equivalently, multipole coefficients) between two of the corner designs?
  (0.2917,0.2083),(0.2917,0.2917)  (0.2917,0.25)?
Not so bad, especially at higher energies, but needs re-tuning iteration

## Slide 47: Horizontal Omni-Magnets
“Arbitrarily” configurable field By(x)
  Are their own correctors
Try flat 80cm-wide poles, 8cm full gap, with large embedded pole face windings
  20 windings each top and bottom
  2x4cm cross-section, 4cm horizontal pitch
Main 8x16cm H-magnet-style coil for dipole
Pole face windings return on nearest side
  Produce quadrupole if all have equal current
Round ones were described in my IPAC’13 paper

## Slide 48: Main Dipole Coil @ 5A/mm2
Roughly flat field, about 0.4T
2D designs in Poisson/SuperFish
(plots are in cm, Gauss)

## Slide 49: Single Pole Winding @ 5A/mm2
Produces ~0.1T field plateau to return side

## Slide 50: Qx=0.2083, Qy=0.2917 F magnet
0.613T peak field, “hardest case”

## Slide 51: Qx=0.2083, Qy=0.2917 F magnet
Winding current densities were found by SVD / least squares

## Slide 52: Qx=0.2917, Qy=0.2917 D magnet
0.404T peak field, only requires small part of aperture

## Slide 53: Qx=0.2917, Qy=0.2917 D magnet
SVD only activates coils in the region they have an effect

## Slide 54: Magnets Summary Table
Cell Qx | Cell Qy | Magnet name | Max current density (A/mm2) | Max By error (G) | Max field (T) | Orbit range (m)
0.2083 | 0.2083 | F | 2.090 | 2.25 | 0.5201 | 0.641
 |  | D | 3.111 | 1.98 | 0.3921 | 0.460
0.2083 | 0.2917 | F | 3.056 | 3.31 | 0.6126 | 0.688
 |  | D | 2.955 | 2.31 | 0.4260 | 0.471
0.2917 | 0.2083 | F | 3.073 | 2.23 | 0.4418 | 0.389
 |  | D | 3.189 | 1.89 | 0.3370 | 0.247
0.2917 | 0.2917 | F | 3.480 | 2.50 | 0.3883 | 0.381
 |  | D | 3.944 | 2.58 | 0.4035 | 0.224
0.216 | 0.213 | F | 2.372 | 1.88 | 0.4224 | 0.573
 |  | D | 4.256 | 2.32 | 0.4270 | 0.400

## Slide 55: Quick Dynamic Aperture Test
Started 8 particles at (x,y) = (±2.5/0,±2.5/0)cm
  Cell Qx=0.216, Qy=0.213 design
Tracked in ring for 60ms, all survived
Some particles at 3cm did not
Can use e = xinj2/binj to get apertures
Energy | Cells (in 60ms) | Turns | ex (geom.) | ey (geom.)
3MeV | 697 | 58 | 471 | 428 mm.mrad
6MeV | 954 | 79 | 408 | 374 mm.mrad
12MeV | 1266 | 105 | 381 | 365 mm.mrad

## Slide 56: Dynamic Aperture Notes
FETS-FFA request was 100mm.mrad normalised at 3MeV (bg=0.080)
  = 1250 mm.mrad geometric
by = 1.46m at 3MeV injection point
  Implies y = ±4.27cm (or ±5.31cm at maximum by)
  Larger than the magnet aperture (±4cm)!
This DA is not the painted beam size, it is a stability margin (all beam ≤ 250 mm.mrad)

## Slide 57: Analytic Method for 2 Thin Lenses
Assume small-angle (paraxial) approximation
  Weak focussing and edge focussing are negligible
  Not valid for the FETS-FFA example!
Fixed cell tunes with two lenses implies constant normalised focussing strength with momentum in each lens
  Qx, Qy are two constraints on linear optics and the focussing strengths (gradients) are two variables
For the case with 3 lenses and 2 constraints, see my IPAC’11 paper

## Slide 58: Function Variable Elimination
xF(p)
xD(p)
x’1(p)
x’2(p)
qF(p)
qD(p)
bF(x)
bD(x)
Four constraints from linear tracking at each energy (and closure)
Two constraints from constant normalised gradient assumption
...becomes the differential equation
Two constraints from rigidity relation
Normalised b = q(Bℓ)
Ends at qcell, Dxcell (constants)
d1
d2
F
D

## Slide 59: Change to Geometric Variables
Rigidity: pq(p) = b(x(p))
Constant norm. gradient: b’(x(p)) = kp
dRigidity/dp: q(p) + pq’(p) = b’(x(p))x’(p)
Therefore: q(p) + pq’(p) = kpx’(p) [= kDx(p)]
x dependency and b have been eliminated
The rest is easy
  Linear algebra
  Solving a 1-variable differential equation
See https://stephenbrooks.org/ap/report/2022-3/fixedtune2.pdf

## Slide 60: Two Thin Lenses Result
The general solution:
C=0 corresponds to straight scaling FFA
  Constant angle, logarithmic orbit position
Range of other solution families are possible
gives field profile implicitly
(arbitrary choice of x origin)

## Slide 61: Example
C values:
Approximately the FETS-FFA parameters but not the same machine (not paraxial)
Negative values of C here produce fields that cross zero like non-scaling FFAs

## Slide 62: Conclusion
Found fixed tune lattices to ~0.001 in cell tune
  And automated: eliminated most hand-tuning!
Dynamic aperture low by factor 2-3× in e
  Not small by most standards: >400mm.mrad
  Haven’t explored the full tune plane yet
Magnetic efficiency about twice as good as (spiral) scaling FFA, magnets buildable
  ~0.6T max field on beam vs. ~1.2T
Very interesting machine type

## Slide 63: Future Work
Dynamic aperture search of tune plane
  And maybe other parameters like fringe length or magnet edge angles (rect vs. sector etc.)
Try expressing transverse field as something better-conditioned than polynomials
  Like Fourier series, or Chebyshev polynomials
Maybe investigate low-level tune “noise” in Muon1 tracking/optics algorithm
  Could also be field model nonsmoothness