## Slide 1: Rotating Coil Results Summary
A
A
B
B
A = Shimmed with 64x 63mil wires, to arbitrary quad strength
B = Shimmed with 64x 80mil wires, to approximate final strength up to temperature variations in bldg. 902 hall
B

## Slide 2
This excludes quadrupole drift from the currently-uncontrolled temperature.
Difference from previous chart comes from taking the max. error on a line rather than 2*RMS on a circle.

## Slide 3: Magnet Development Tasks
Task | Status
Multipole reduction | Done for 3 QF magnets. Showed 8-10x improvement but still marginal to spec for 2 BD magnets. May not be possible for other BDs due to construction errors.
Temperature control & field stability | This test is planned as soon as John Cintorino has time to do it.
Strength tuning | Some rough control demonstrated but needs water cooling to do well, this test is planned.
Reassembly repeatability | Did not work well due to construction errors in QF4. Maybe try a better-made magnet?
Effect of window-frame | This test can be done in future. Just need to ensure window-frame does not move multipoles outside tuning range.

## Slide 4: BACKUP

## Slide 5: FFAG preproduction magnet tests and production testing plan
“Preproduction” = “First girder”

## Slide 6: Magnets Made So Far
“ATF/PoP” magnets (6x QF, 6x BD, half length)
  Based on old 250MeV CBETA design
  Slightly smaller bores, unequal for QF and BD
“First girder” magnets (4x QF, 4x BD, 2 layers)
  Based on 150MeV “v3” design, few % different
  Full aluminium surround (stronger)
  Splittable into halves
    Requires use of glue and metal alignment pins
  Water cooling
FFAG’16 talk
IPAC’17 paper
Backup slides

## Slide 7: Example First Girder QF

## Slide 8: Bad News First
During assembly, after gluing, but before measurement, ~1mm size errors were observed by eye
  Compare ±0.25mm (0.01”) position tolerance
Strategic decision to meet April 30th milestone deadline with “sub-optimal” assembly
  Instead of halting production, finding and fixing the flaw before continuing & being 1-2 weeks late
    In production I hope we choose the latter path

## Slide 9: “Panel Gap” Problem (QF4)

## Slide 10: “Panel Gap” Problem (BD2)
Steve Bellavia’s photograph

## Slide 11: “Panel Gap” Problem (BD2 zoom)

## Slide 12: Gap Size from Photography
1.42mm

## Slide 13: Fault Tree Investigated
Aluminium halves not touching
  Jammed on alignment pins? No (can happen but fixes itself under force)
  Dirt or glue between halves? Some, but gap remained after careful cleaning of surfaces
  Permanent magnets protruding? Yes: halves are resting on corners/sides of end magnets
    Gradual creep of epoxy glue? No, as far as we can tell
    Misplacement during gluing? Most likely explanation

## Slide 14: Misplacement During Gluing
What should happen
What did happen (current theory)
Blue layer is ~1/8” Teflon added to prevent glue adhering to the ½” aluminium plate below
With the two end blocks allowed to move outwards, the rest of the blocks are no longer packed tightly (with the brass shims) so position errors can be everywhere

## Slide 15: Rotating Coil Results Summary

## Slide 16: Rotating Coil Results Summary
A
A
B
B
A = Shimmed with 64x 63mil wires, to arbitrary quad strength
B = Shimmed with 64x 80mil wires, to approximate final strength up to temperature variations in bldg. 902 hall
B

## Slide 17: “Panel Gap” Problem (QF4)
This is why reassembly relying on the alignment pins didn’t work well

## Slide 18: Quadrupole Strength Errors
A
A
B
B
A = Shimmed with 64x 63mil wires, to arbitrary quad strength
B = Shimmed with 64x 80mil wires, to approximate final strength up to temperature variations in bldg. 902 hall
B

## Slide 19: Conclusions (whole magnets)
Iron wire tuning works well as long as the original magnet is “within range”
Reassembly won’t be repeatable until we have better construction
To-do:
  Water cooling test (time constant, ripple)
  Absolute strength tuning (requires water cooling)
  Possible rebuild/modification of a magnet

## Slide 20: Helmholtz Tests
All the individual PM blocks for the first girder were measured in a 3-axis Helmholtz coil
32 blocks per magnet * 12 magnets = 384 blocks (we had spares)
Thanks to Steve Trabocchi
3D printed holders for each block type

## Slide 21: Magnetisation Angle Errors

## Slide 22: Magnetisation Strength

## Slide 23: Flux Vector (BD pieces)

## Slide 24: Expected vs. Spec
Block Type | Quantity | Average | St.Dev. | RMS | Spec.
QF all | Strength (T) | 1.18275 | 0.00508 (0.429%) | n/a | Average 1.17—1.22 s ≤ 0.6%
QF all | Angle error (°) | -0.70164 | 0.61195 | 0.92996 | RMS ≤ 1%
BD (P1-12) | Strength (T) | 1.18495 | 0.00553 (0.466%) | n/a | Average 1.17—1.22 s ≤ 0.6%
BD (P1-12) | Angle error (°) | 0.18826 | 0.47064 | 0.50537 | RMS ≤ 1%
BD all | Strength (T) | 1.18109 | 0.00892 (0.756%) | n/a | Relax on smallest blocks
BD all | Angle error (°) | -0.24124 | 0.96941 | 0.99652 | 
Spec. values from CBETA note 10

## Slide 25: Conclusions (Helmholtz)
Results were good
  Mean and standard deviation as expected
  In both block strength and magnetisation angle
  No weird long tails: normal distribution or better
Iron wire tuning should work fine
The part we messed up in these magnets is construction, which is the part BNL is not going to do (the vendor will)

## Slide 26: Production Testing Plan
Helmholtz measurements of at least ~10% of individual blocks
  Line item cost options: 100% done by vendor, or 10% done by vendor, or (fallback) 10% at BNL
Rotating coil measurement of whole magnets
  At least 2 measurements: before and after iron wire multipole tuning
  Needs water cooling circuit attached + corrector
  Most likely done at BNL as we have 10-5 coil
Determines initial thickness of brass inter-block shims
Can also course-correct brass shim thickness

## Slide 27: BACKUP
Bellavia’s photo processing
First girder magnets design and ATF magnets results

## Slide 28: Bellavia photo processing (1/3)

## Slide 29: Bellavia photo processing (2/3)

## Slide 30: Bellavia photo processing (3/3)

## Slide 31: FFAG Magnet Requirements & Specifications
Including results of Halbach proof-of-principle magnets and iron wire shimming

## Slide 32: Magnetic Design
Magnet name | Length (mm) | Dipole at centre (T) | Quadrupole (T/m) | Radius to magnet pieces (mm) | Radius incl. shim holder (mm)
QF | 133.3 | 0 | -11.5 | 42.5 | 39.4
BD | 121.7 | -0.311 | 11 | 42.5 | 39.4
BP1.5
CBETA note #1
DR3.1.2
DR2.6.1
DR = Design Report
BP = Baseline Parameters

## Slide 33: Baseline Parameter Sheet #1.5

## Slide 34: Beam-Quadrupole Errors Path
Quadrupole errors on beams:
±0.2% skew from roll errors
±0.12% from temperature
±0.1% from strength errors
Magnet position errors
Magnet roll errors
±2mrad = ±0.11°
Temperature variation magnet blocks
±1.1K
Magnet multipole errors
Magnet strength errors
±0.1%
Dotted lines only have a tiny contribution via sextupole feed-down
Conversion factor: 1 (skew)
Conversion factor: 1
Conversion factor:
NdFeB temperature coefficient = -1.1e-3/K
BP1.5
BP1.5

## Slide 35: Quadrupole Corrector not Powered
±0.22% quadrupole error and ±0.2% skew only produces a small beta ripple within the beam
  Acceptable without correction
Quad corrector is not powered in baseline
  Dipole-on-beam effects that feed down from the quadrupole errors can be corrected as “random” dipole errors as shown next
±2% quad corrector is included as an option if there is no water cooling (±18K range)

## Slide 36: Beam-Dipole Errors Path
Dipole errors on beams (T)
Magnet position errors
Magnet roll errors
Temperature variation magnet blocks
Air temperature variation
Water temperature variation
Magnet multipole errors
Magnet strength errors
Required dipole corrector field
Block magnetisation angle errors
Block magnetisation strength errors
Assembly block placement errors
Iron wire shimming
Each box is a value (or range size)
Each arrow is a conversion factor
Corrector windings heat source
Radiation damage

## Slide 37: Correction Simulation Result
Simulations done by Chris Mayes using BMAD
2.1 Gauss.cm/(um offset), at >=11T/m
2.1 Gauss.cm/0.11 Gauss = 19.1 cm conversion factor to get integrated field
DR2.13

## Slide 38: Required Correction Strength
This uses the current CBETA spec for the dipole corrector
Corrector is 12.01cm long
  Conversion factor is 1.59*L
Can rewire quadrupole coils as dipole for extra strength
Dipole errors on beams
±39.6 Gauss
Required dipole corrector field
±757 Gauss.cm
Conversion factor:
19.1cm
757 G.cm
757 G.cm
+quad
1342 G.cm
BP1.5

## Slide 39: Position Errors
Dipole errors on beams:
±23.0G from position errors
±16.3G from roll errors
Magnet position errors
±200um
Magnet roll errors
±2mrad = ±0.11°
= ±200um at R=10cm
Conversion factor: gradient
11.5 T/m worst case
Conversion factor:
central dipole field + gradient * beam radius * 2
0.311 + 11*0.023*2 = 0.817 T worst case
BP1.5
BP1.5

## Slide 40: Temperature
Dipole errors on beams:
±6.8G
Temperature variation magnet blocks:
±0.99K from water
±0.08K from air
±0.03K from corrector
Total ±1.1K
Air temperature variation
15-30C = ±7.5K
Water temperature variation
19-21C = ±1K
Conversion factor:
relative temperature coefficient of NdFeB material * dipole field on beam
-1.1e-3/K * 0.564 T (worst case)
= 6.2 Gauss/K
Corrector windings heat source
5.37W dipole + 31.53W quadrupole
= 36.9W
Conversion factor: 0.99
Conversion factor:
0.7mK/W
Conversion factor: 0.01
BP1.5
BP1.5
BP1.5
Above factors are from George Mahler’s thermal simulations

## Slide 41
12 proof-of-principle magnets (6 QF, 6 BD) have been built as part of CBETA R&D.
Iron wire shimming has been done on 3 QFs and 6 BDs with good results.
PoP QF
PoP BD
Iron wire shims
PoP magnet series

## Slide 42: Proof-of-Principle vs. CBETA
Parameter | CBETA QF | PoP QF | CBETA BD | PoP BD | Units
Length | 133.30 | 57.44 | 121.70 | 61.86 | mm
Gradient | -11.500 | -23.624 | 11.000 | 19.119 | T/m
Dipole at centre | 0 | 0 | -0.31100 | -0.37679 | T
Good field radius | 23.00 | 20.20 | 23.00 | 13.70 | mm
Blocks inner radius | 42.50 | 37.20 | 42.50 | 30.70 | mm
Blocks outer radius | 55.76 | 62.45 | 72.89 | 59.43 | mm
Max field in good field region | 0.265 | 0.477 | (-)0.564 | (-)0.639 | T
Max field at “pole tip” | 0.489 | 0.879 | (-)0.779 | (-)0.964 | T
NdFeB material grade | N35EH | N35SH | N35EH | N35SH | 
DR3.1.4

## Slide 43: Magnet Quality
Dipole errors on beams:
±2.6G from strength errors
±2.6G from multipole errors
Magnet multipole errors
~10-3 (10 units) at R=25mm
Magnet strength errors
±0.1%
Block magnetisation angle errors
Initial factory guarantee: ±5°
New factory guarantee: ±3° and 80% within ±2°
Observed RMS: 1.14°
Block magnetisation strength errors
Factory guarantee: ±2.1%
NB: mean could be shifted in this range even if RMS is small
Observed RMS: <2%
Assembly block placement errors
Estimated: 0.25mm
New aluminium frame gives <0.13mm flex
Iron wire shimming
Iron wire shimming will reduce all initial multipole errors to about the same level unless they exceed a threshold
Detailed multipole tolerances on following slides but for previous lattice, so may change slightly
Conversion factor:
gradient * beam radius
11.5*0.023 = 0.2645 T worst case
BP1.5

## Slide 44: Iron Wire Shimming Improvement
Factor of ~4 reduction regardless of starting value
2nd iteration improved further in previous R&D:
29.62  6.69  1.94
DR3.1.4

## Slide 45: Multipole Error Simulations
Effectively, the multipoles just create another combination of dipole errors on each beam
Chris Mayes did the study in BMAD
Created tables of recommended multipoles
Dipole errors on beams (T)
Magnet multipole errors
Required dipole corrector field
Beam dynamics
DR2.14.6

## Slide 46: Multipole Tolerances
DR2.14.6

## Slide 47: Iron Wire Shimmed QF Results
Below ~10 units at R=25mm, equivalent to 0.1% of main strength
DR3.1.4

## Slide 48: Quadrupole Variation
Larger than 0.1% strength variation
What is different?
Not a ratio: absolute strength, affected by temperature coeff.
DR3.1.4

## Slide 49: Improving Strength Consistency
The proof-of-principle magnets had no temperature control during measurement
  Expect 0.11% per K difference between initial unshimmed measurement and re-measurement
  Largest is equivalent to 6.3K difference (902 Hall)
  For first girder magnets will run water circuit
Could also use “reference” magnet mounted backwards on same rotating coil
  Iron wire shim total quadrupole to zero

## Slide 50: Improving Average Strength Offset
Magnets are designed for low end of stated magnetisation range (“too much” material)
Weakened by moving blocks radially outwards by up to 0.8mm with inter-block Al shims
  Compensates whole-batch average strength
Magnet | Radial block displacement (mm) | Dipole (T) | Quadrupole (T/m) | Change in dipole (%) | Change in quadrupole (%)
QF | 0 | 0 | -11.5 |  | 
QF | 1 | 0 | -10.8414 | - | -5.727%
BD | 0 | -0.311 | 11 |  | 
BD | 1 | -0.29991 | 10.4222 | -3.566% | -5.253%
CBETA note #1

## Slide 51: Radiation Tolerance
Lifetime dose limit of 1kGy = 100krad
  10 rad/h at beam pipe average over 10000 hours
  Easily monitorable with dosimeters
Dipole errors on beams:
0.36G
Magnet multipole errors
1.3 units at R=25mm
Magnet strength errors
-1.3 × 10-4
Radiation damage
1kGy = 100krad
Conversion factors:
1% / 74.5kGy
= 0.134/MGy
Conversion factor:
0.2645 T worst case
BP1.5

## Slide 52: Recap: Beam-Dipole Errors Path
Dipole errors on beams
Total ±52.0G
Magnet position errors
±200um
Magnet roll errors
±2mrad = ±0.11°
Temperature variation magnet blocks
Total ±1.1K
Air temperature variation
±7.5K
Water temperature variation
±1K
Magnet multipole errors
~10 units at R=25mm
Magnet strength errors
±0.1%
Required dipole corrector field
±993 Gauss.cm
Block magnetisation angle errors
±3° and 80% within ±2°
Block magnetisation strength errors
±2.1%
Assembly block placement errors
Estimated: 0.25mm
Iron wire shimming
Each box is a value (or range size)
Each arrow is a conversion factor
Corrector heat source
36.9W
Radiation damage
1kGy
0.99
0.7mK/W
0.01
0.134/MGy
6.2 G/K
0.2645 T
0.2645 T
11.5 T/m
0.817 T
19.1cm
BP1.5
BP1.5
BP1.5
BP1.5
BP1.5
BP1.5
BP1.5

## Slide 53: Dipole Error Contributions
757 G.cm
1342 G.cm
BP1.5

## Slide 54: Conclusion
All sources of field error on the beam have been estimated
Corrector strength required per unit field error is known from simulations
These have been put into a corrector budget
  We can correct the worst case with the hardware that is currently in our costing

## Slide 55: BACKUP

## Slide 56: Correction Simulations
Simulations done by Chris Mayes using BMAD
Dipole errors on beams (T)
Required dipole corrector field
Beam dynamics
(N=100)
DR2.13

## Slide 57: Orbit Correction of 4 Beams
DR2.13

## Slide 58: Herding Cats
Before correction
After correction

## Slide 59: Herding Cats
After correction
Same X, Y correctors on all 4 beams
Beams have different phase advances
  Over >4 cells correctors become linearly independent on all beams
Max corrector 1.99x max error

## Slide 60: Dipole Corrector Boost
Configuration | Coils | Amps | Volts | Watts | Ohms | Turns/coil | Amp.turns/coil | Dipole field (G.cm)
Power supply |  | 3 | 8 | 24 | 2.667 |  |  | 
Dipole corrector | 2 | 2.08 | 2.58 | 5.366 | 1.240 | 390.4 | 812.1 | 757
Quad corrector | 4 | 5.26 | 5.995 | 31.53 | 1.140 | 301.9 | 1588 | 1480
Half of quad corrector | 2 | 5.26 | 2.998 | 15.77 | 0.5699 | 301.9 | 1588 | 1480
Dipole + half of quad Everything powered at 2.08A | 4 | 2.08 | 3.765 | 7.832 | 1.810 | 692.3 | 1440 | 1342
BP1.5
BP1.5
BP1.7 & quote

## Slide 61: Mechanical Simulation (G. Mahler)
Maximum 0.005” deflection under magnet block forces
= 0.127mm << block random alignment errors in prototypes  shimming OK

## Slide 62: Thermal Simulations (G. Mahler)
60F water temp
5W/m2 convection coefficient,
100W top, 50W each side
Maximum 0.5F variation at blocks
0.5F/400W = 0.7mK/W

## Slide 63: Air Temperature Effect
Air +10F
Almost no change at blocks
<0.1F

## Slide 64: Proof-of principle Design and Parameters
Parameter | “QF” magnet | “BD” magnet | Units
Length | 57.44 | 61.86 | mm
Dipole By(x=0) | 0 | -0.37679 | T
Quadrupole dBy/dx | -23.624 | 19.119 | T/m
Inner radius (magnet pieces) | 37.20 | 30.70 | mm
(shim holder) | 34.70 | 27.60 | mm
“Pole-tip” field (magnet pieces) | 0.879 | (-)0.964 | T
Outer radius (magnet pieces) | 62.45 | 59.43 | mm
(tubular support) | 76.2 | 76.2 | mm
QF
BD
Based on Dec 2015 lattice, when pieces were ordered (Emax=250MeV)
DR3.1.4

## Slide 65: Unshimmed Magnet Errors
Magnet | Reference radius | RMS of FOM (units) where FOM = ∑multipoles2 | Simulated: X,Y ±0.25mm Strength ±2.1% Mag. angle ±3°
QF | 10mm | 10.2434 | 28.6128
QF | 25mm | 46.9252 | 88.8997
BD | 10mm | 28.8490 | 37.9671
BD | 25mm | Not measured | 116.947
As expected, measured spread is slightly better than the “worst case” of uniform random errors filling the whole ranges

## Slide 66: Temperature of 902 Hall
6K
6K
Variations consistent with observed ~0.7% maximum magnet strength variation
(equivalent to 6.3K change)

## Slide 67: Radiation Resistance
Temnykh 2008 formula for 1% field loss @ RT
  Dose1%[Mrad] = 10^(-2.68+Tdemag(°C)/41.4K)
    Where Tdemag is temperature where material grade would demagnetise given its position in the magnet
The N35EH grade is rated at Tdemag = 200°C when isolated (self-generated H field only)
  Halbach magnet contains up to 1.3T of reverse flux (antiparallel H field), so reduce this to 147°C
Dose1% = 7.45Mrad = 74.5kGy = 74.5 kJ/kg

## Slide 68: Temperature Dependent B-H Curve
NB: this is N35SH shown for example
kOe is kG with a factor of m0
100K change produces 11.9kOe change, i.e. 1.19T
Rated temperature for N35SH is 150°C, at -0.67T X-axis intercept
200-100*(1.3-0.67)/1.19
= 147°C
-1.3T

## Slide 69: Multipole Errors within Beams
Negligible, beams are small
  10 units of sextupole at R=25mm becomes 0.4 units or less at R=1mm, regardless of beam centre
Swamped by dynamical nonlinearities
  Dynamic aperture is observed in simulations, but beam will only hit it if its centroid is displaced by several mm